## Slide 1: Getting Data in the Lab
LabVIEW, Data Acquisition,
and Instrument Control
Dr. Jason D. Sommerville
President, UPVI, LLC.
jason.sommerville@upvi.net
www.upvi.net
Copyright © 2015 UPVI, LLC. Redistribution is prohibited.

## Slide 2: Getting Data in the Lab
Part 2: Data Acquisition
Copyright © 2015 UPVI, LLC. Redistribution is prohibited.

## Slide 3
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics

## Slide 4: Data Acquisition
Data Acquisition or DAQ is the process of converting real-world measurements into digital computer data
  Signal measurement
  Signal digitization
  Moving the data to computer memory and/or disk
DAQ may also encompass signal generation
  Waveform generation
  Triggering
  Control systems
All the components must be compatible with each other!
  But we have to break it down in the lecture and handle one thing at a time
In 2016, if you’re writing data down by hand, you’re probably doing it wrong

## Slide 5: Types of DAQ: Signal Acquisition
Convert analog voltage signals into digital data
Single-shot or timed measurements
Analog Input
Digital Input
Counter/Timers
Use an analog-to-digital converter
Examples: temperature, sounds, voltage
Single-shot or timed
Single bit or bus
Example: proximity sensor, photointerrupter
Count events
Time events
Examples: encoders, pulse train timing, delay timing, delay generation
101110110100 101110100101 011001100101
110011000101 011100101010 110110001000
1 or 0
Buffer
555
4, 5, 6, …
10 ms
10 ms
Counter/Timer

## Slide 6: Types of DAQ: Generation
Convert digital data into analog waveforms
Single-shot or timed generation
Analog Output
Digital Output
Counter/Timers
Use a digital-to-analog converter
Examples: sound, system excitation, motor control
Single-shot or timed
Single bit or bus
Generating a slow trigger, controlling an LED, controlling a bus
Control timing
Examples: Pulse train generation, delay generation
1 or 0
Buffer
555
4, 5, 6, …
10 ms
10 ms
Counter/Timer
101110110100 101110100101 011001100101
110011000101 011100101010 110110001000

## Slide 7
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Transducer
Thermocouple

## Slide 8: Transducers
Convert real world data into something that can be measured by a data acquisition device
  Voltage
  Current
  Resistance
Also called sensors

## Slide 9: Transducer Menagerie
Temperature | Pressure | Flow
Thermocouple Thermistor RTD IC Temp Sensor IR Sensor Fiber Bragg Grating (FBG) | Piezo-resistive strain gauge Capacitive Electromagnetic Piezoelectric Microphone FBG Ion gauge | Differential Pressure Pitot tube Calorimetric Turbine Electromagnetic Ultrasonic
Light | Position/Velocity | Strain, weight, etc
Photodiode Phototransistor Photomultiplier Single-photon avalanche diode CCD/CMOS array | Optical Encoders (A,L) LVDT (L) Resolver (A) Potentiometer (A) String Pot (L) Radar (L) Laser interferometry (L) | Strain gauge Load cell Piezoelectrics

## Slide 10: How to chose a transducer
What type of measurement do you need to make?
Calibration Curve
  Range
    Input range (25 – 300 ºC, 1 – 100 mN, etc.)
    Output range (0 to 10V, -5 to 5 V, 4 to 20 mA, etc.)
  Sensitivity—How well can change be detected?
How good is the measurement?
  Accuracy (±1 ºC, ±0.1 mN, etc.)
  Precision (repeatability)
  Linearity—can be handled in software
  Drift—Change in calibration over time
  Hysteresis—is it the same going up as coming back down?
Timing
  Response time
  Bandwidth (1 Hz, 100 kHz – 1 MHz, etc.)
Other considerations
  Analog or digital
    Analog: voltage, current, impedance, etc.
  What signal conditioning or excitation/power does it require?
  How will I connect it to my DAQ system?
  Environmental isolation
    Voltage, temperature, noise, vacuum, water, etc.
  Ruggedness
  Cost
A: All this info should be on the transducer’s datasheet or specification sheet
Q: Where do I find this stuff?
(which should be available from the manufacturer’s website)

## Slide 11: Range
Range specifies the maximum and minimum values a transducer can accurately handle
Input range (or just “range”) is the values of the real world signals that can be accurately measured
  0 – 100 degrees C
  0 – 10 mN
  -500 – 500 mm
  Should cover the expected range of your physical system!
Output range is the corresponding range of output:
  0 – 10 V
  -2 – 5 V
  4 – 20 mA
  0 – 100 kΩ
  Must work with your DAQ system!

## Slide 12: Sensitivity
How much does the output change for a given change in input?
  That is, the slope of the calibration curve
Units of output unit/physical unit
  V / deg C
  V / N
  mV / mrad
  mA / (L/s)
Tradeoff between sensitivity and range
Sensor may not be equally sensitive across its full range
  I.e., calibration curve may not be linear

## Slide 13: Calibration Curve
Relates the physical input signal to the electrical output signal
A linear function is used to relate input (x) to output (y).
Linear
Mapped ranges
Different way of specifying a linear calibration curve

## Slide 14: Calibration Curve
Relates the physical input signal to the electrical output signal
A general function is used to relate input (x) to output (y), e.g.
Functional, e.g polynomial
Table, i.e. piecewise linear
Linearly interpolate between measured points.
Useful for measured calibration curves

## Slide 15: Example
Thermocouples measure the difference in temperature between the hot junction and the cold junction
There are different types of thermocouples for different temperature ranges
More reading: http://www.ni.com/white-paper/4218/en/

## Slide 16: Range: Choosing a thermocouple
http://www.temprel.com/support/how-to-choose-thermocouple.aspx

## Slide 17: Linearity and Sensitivity: Ex: Choosing a thermocouple

## Slide 18: How well does it measure?

## Slide 19: Response/Rise/Fall Time
How fast can the sensor respond to changes in the measured property?
Hall sensor
Rise Time
Fall Time

## Slide 20: Bandwidth
Usually another way of specifying the response time
Note: For response time given as 10% to 90% rise/fall

## Slide 21: Drift
The calibration of a sensor may change over “long” periods of time
  Temperature effects
  Humidity effects
  Aging effects
What is long? Depends on your measurement.
Source: http://www.vectornav.com/support/library/calibration

## Slide 22: Hysteresis
Image: http://www.ni.com/white-paper/14860/en/#toc8
Is it the same going up as coming down?

## Slide 23: Reading a Spec Sheet
Show examples of ways in which specifications are given.

## Slide 24: Analog vs. Digital
Analog
Voltage, current, impedance, etc.
Requires DAQ device
Typically faster
  Sensor response time limits update speed
Typically cheaper
Transmission lines susceptible to noise
More difficult/expensive to isolate
Digital
Communicates via GPIB, RS-232, USB, Ethernet, etc.
    May require interface card, or may be built into system
    Typically slower and more expensive
      Comm limits update speed
    Typically more expensive
    Less susceptible to noise
    Easier to isolate
Nearly all transducers are analog internally
However, some are packaged with dedicated ADCs, signal conditioners and may provide digital communication outputs

## Slide 25: Example: Gas Flow Meter
Analog: Omega FMA-3103
No internal calibration curve
Analog output only
  Output is linearized
  => internal signal conditioning
Update rate of 0.5 Hz
Cost: $375
Digital: Omega FMA-4102
Internal, programmable gas calibration curves
Interface: RS-232 or RS-485
    Unspecified, probably about 2 Hz
    Cost: $730
http://www.omega.com/pptst/FMA3100_3200_3300.html
http://www.omega.com/pptst/FMA4100_4300.html

## Slide 26: Transducers: Other Considerations
What signal conditioning or excitation/power does it require?
How will I connect it to my DAQ system?
  Cabling
  Isolation (see next bullet)
Environmental isolation
  Voltage, temperature, noise, vacuum, water, etc.
Ruggedness
Cost

## Slide 27: How to chose a transducer
What type of measurement do you need to make?
Calibration Curve
  Range
    Input range (25 – 300 ºC, 1 – 100 mN, etc.)
    Output range (0 to 10V, -5 to 5 V, 4 to 20 mA, etc.)
  Sensitivity—How well can change be detected?
How good is the measurement?
  Accuracy (±1 ºC, ±0.1 mN, etc.)
  Precision (repeatability)
  Linearity—can be handled in software
  Drift—Change in calibration over time
  Hysteresis—is it the same going up as coming back down?
Timing
  Response time
  Bandwidth (1 Hz, 100 kHz – 1 MHz, etc.)
Other considerations
  Analog or digital
    Analog: voltage, current, impedance, etc.
  What signal conditioning or excitation/power does it require?
  How will I connect it to my DAQ system?
  Environmental isolation
    Voltage, temperature, noise, vacuum, water, etc.
  Ruggedness
  Cost

## Slide 28
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Signal Conditioner
Amplifier, Low pass filter
Transducer
Thermocouple

## Slide 29: Signal Conditioning
Modify the transducer output so that it can be accurately measured by the DAQ system
Some basic signal conditioning is often included with some transducers and some DAQ systems
  X -> Voltage conversion
  Amplification
  Anti-aliasing filters
Amplification/Attenuation
  Small signals (typ. <100 mV) should be amplified to a range readable by the ADC
  Large signals (>5 – 10 V) voltage signals need to be attenuated
Filtering
  Noise: use low-pass or notch filters to remove known noise sources
  Anti-aliasing filter: a low-pass filter designed to prevent high-frequency components from being digitized and aliased by the DAQ system

## Slide 30: Signal Conditioning (cont.)
Voltage Isolation
  Allows signals at high potential differences to be safely connected to a DAQ system
Linearization
  Corrects for non-linearities in the transducer
  In modern systems this is usually done in software
Transconductance, etc.
  Convert current signals to voltage signals
  Other conversions possible as well (i.e. resistance to voltage, pulse-train to voltage, etc.)
Excitation
  Some sensors, e.g. RTDs and strain gauges, require a power source to excite them
  Not strictly signal conditioning

## Slide 31: Signal Conditioning—Hardware
Amplifiers and filters often build into DAQ hardware
Rack and rail systems
  NI SCXI
  Omega DRG or OM7
Custom solutions
  Build it yourself
NI SCXI
Omega DRG Module

## Slide 32
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ

## Slide 33: How to choose DAQ hardware
What type of signals do you need to acquire/generate?
How many channels of each type?
After external conditioning, what will the polarity and range of your signal be?
DI/SE?
What resolution do you need across that range?
What bandwidth do you need?
At what rate do you need to acquire/generate the data?
  Note that bandwidth ≠ sampling frequency, but related
  How precise are you timing requirements? Can you use software timing?
  How much time-skew is acceptable?
Memory depth if signals cannot be transferred to system memory
Connectors
Form factor
Computer interface (e.g. PCIe, PXI, USB, etc.)
Software interface
Front panel controls
Cost

## Slide 34: Analog Input Digitization
Analog signal: 3 Hz Sine
Sampling Period: 50 ms (20 Hz)
Digitized Samples
Sample Index | Time (s) | Voltage (V)
0 | 0.00 | 1.7101
1 | 0.05 | 4.80631
2 | 0.10 | 3.94005
3 | 0.15 | -0.174497
4 | 0.20 | -4.14519
5 | 0.25 | -4.69846
6 | 0.30 | -1.37819
7 | 0.35 | 3.07831
8 | 0.40 | 4.99695
9 | 0.45 | 2.79596
10 | 0.50 | -1.7101
11 | 0.55 | -4.80631
12 | 0.60 | -3.94005
13 | 0.65 | 0.174497
14 | 0.70 | 4.14519
15 | 0.75 | 4.69846
16 | 0.80 | 1.37819
17 | 0.85 | -3.07831
18 | 0.90 | -4.99695
19 | 0.95 | -2.79596

## Slide 35: DAQ: Analog Input
Central component is the Analog-to-Digital (ADC) converter
  Digitizes the analog signal and making a digital measurement available for storage in memory
  ADC must be supported by amplifiers, clocks, multiplexers, and other items
Typical DAQ-AI system has the following parameters:
  Differential vs. Single-ended inputs
  Voltage range and polarity (bipolar or unipolar)
  Gain settings
  ADC resolution
  Sampling speed and bandwidth
  Multi-channel multiplexing

## Slide 36: Analog Input
Differential Input (DI)

## Slide 37: Analog Input
Single-ended Input (SE)

## Slide 38: Analog Input: DI vs. SE

## Slide 39: Analog Input: DI vs. SE
Differential Input
Used with balanced signal lines
Less susceptible to noise
Less likelihood of ground loops
Lower channel count
Single-Ended
Used with unbalanced signals
Increased susceptibility to noise
Increased likelihood of ground loops
Higher channel count
Voltage difference must be within DAQ range
Neither side must float too far from DAQ ground
Differential Input Caveats

## Slide 40: Analog Input
Analog signal: 3 Hz Sine
Sampling Period: 50 ms (Rate is 20 Hz)
Digitized Samples
Bi-polar, Range: ±5 V
Resolution: 16 bits across the range

## Slide 41: ADC Range and Polarity
Range: the maximum and minimum voltage that may be read
Unipolar: only read signals >= 0 V (or only <= 0, sometimes)
Bipolar: may read positive and negative signals
Bi-polar, Range: ±5 V
Most Common Ranges:
±5 V, ±10 V

## Slide 42: Gain Settings
Configurable gain settings allow ADC to make more use of available ADC resolution
Note that noise is also amplified—better to amplify at source, if possible
Signal here must be (e.g.) ±5 V to make full use of ADC resolution
Gain is software-configurable

## Slide 43: Gain Settings
Working range <= signal range
e.g.: ±250 mV
Gain, e.g. 20
ADC range, e.g. ±5 V

## Slide 44: Gain Settings: Specifications
Spec. Ranges
Spec. Gains
Multiple ranges specified
Gain unspecified
Actual ADC range unspecified
Example: NI PCI-6251
  http://www.ni.com/pdf/manuals/375213b.pdf
Multiple gains specified
Only maximum range specified
ADC range must be specified
Example: Keithley KUSB-3101
  (Actually specifies both)
  http://www.keithley.nl/data?asset=50072

## Slide 45: ADC Resolution: 3 bits
Number of bins = 23 = 8
Range = 5V – (-5V) = 10 V

## Slide 46: ADC Resolution: 4 bits
Number of bins = 24 = 16
Range = 5V – (-5V) = 10 V

## Slide 47: ADC Resolution: 8 bits
Number of bins = 28 = 256
Range = 5V – (-5V) = 10 V

## Slide 48: ADC Resolution
The resolution with which a signal can be measured, specified in bits
Resolution in volts is given by ADC Range, Gain setting, and ADC resolution
Note that sometimes the specified “range” will be the product of the smallest gain and range, or all range-gain products will be specified.
Most Common Resolutions:
8, 12, 16, 24
Example:

## Slide 49: Voltage <-> ADC Reading
Example:, 5 – (-5) V range, 16 bit resolution, reading = 8745
Other direction:

## Slide 50: Sampling Speed
The maximum rate that the hardware can make ADC conversions
In samples per second
  E.g. 250 S/s, 350 kS/S, 10 MS/s
Hardware or software controlled?
  Low-end hardware must be clocked by software—very unreliable timing on modern Oses, no better than 100 ms
  Lab-quality software will use hardware timing
What about bandwidth?
  A system may be able to sample above its bandwidth
  Imagine putting a 10 Hz low-pass filter ahead of your 10 MS/s ADC
  Bandwidth is set by the analog characteristics of the electronics
  Bandwidth may be affected by your setup, esp. cabling.
  Usually not a concern until you get to 10s – 100s of MHz

## Slide 51: Undersampling
Undersampled data causes aliasing
The alias looks like a lower frequency signal
Analog signal: 35 Hz Sine
Sampling Period: 50 ms (20 Hz)
Digitized Samples
Alias Period: 200 ms (5 Hz)

## Slide 52: Nyquist-Shannon Sampling Theory
“The sampling frequency must be greater than or equal to twice the bandwidth of the signal sampled to determine the frequency content of the sample.”
If bandwidth starts at 0 Hz (usual), BW is simply the maximum frequency component in our sample.
If you need to determine the non-sinusoidal shape of a signal repeating at X Hz

## Slide 53: Anti-alias Filters
If you know you have frequency components which are not interesting, use an anti-alias filter.
Prevents aliasing by preventing the signal!

## Slide 54: Trade off: Resolution or Speed
For the same priced-system, there is a tradeoff between resolution and sample speed.

## Slide 55: Analog Input: Multiplexing
Sample Clock: Acquire one data point from every channel
Convert Clock: Acquire one data point from any channel
Convert Clock
Sample Clock/N
Sample Clock

## Slide 56: Analog Input: Time Skew
Scan Clock
10 Hz
Sample Clock
100 Hz

## Slide 57: Dealing with Time Skew
Live with it
  Does it really matter for your data?
  Make sure you understand the effect it will have on your data
  By far the most common solution
Sample-and-hold circuitry
  An analog circuit which captures all channels’ voltages simultaneously
  Allows ADC and MUX to cycle through sampled voltages at scan clock rate
Multiple ADCs

## Slide 58: Analog Input: Full Disclosure
Gain circuitry actually comes after the MUXs in most designs.
This explains the tradeoff between DI/SE and channel count

## Slide 59: Some Examples
NI CompactDAQ Platform
USB connectivity or embedded controller
Modules support AI, AO, DIO Counter/timers, some signal conditioning
$800+
NI USB-6008
USB connectivity
12-bit AI 8 SE/4 DI
10 kS/s
2 AO
$189
Meas. Computing PCI-2513
PCI card
16-bit AI 16 SE/8 DI
1 MS/s
DIO, counter/timers
$909
NI PXIe Platform
Embedded controller or PCIe bridge
Modules support everything
$2000+

## Slide 60: How to choose DAQ hardware
What type of signals do you need to acquire/generate?
How many channels of each type?
After external conditioning, what will the polarity and range of your signal be?
DI/SE?
What resolution do you need across that range?
At what rate do you need to acquire/generate the data?
  Note that bandwidth ≠ sampling frequency, but related
  How precise are you timing requirements? Can you use software timing?
  How much time-skew is acceptable?
What bandwidth do you need?
Memory depth if signals cannot be transferred to system memory
Connectors
Form factor
Computer interface (e.g. PCIe, PXI, USB, etc.)
Software interface
Front panel controls
Cost

## Slide 61: Recall: Analog Input
Differential Input (DI)

## Slide 62: Analog Output

## Slide 63: Analog Output Hardware
For the most part, specs are exactly like Analog Input, but “in reverse”
Range: Voltage output range
Polarity: Bipolar or unipolar output
Gain: Typically only the working range is specified.
  Selectable gains/working ranges far less common for output
Resolution: Specifies minimum output step size
Differential/Single Ended: yes, but
  Almost always single ended
  Why? Because we’re now the output. We choose the reference.

## Slide 64: Analog Output

## Slide 65: Analog Output Hardware: Timing
For the most part, specs are exactly like Analog Input
Timing: Hardware or software timed
  Devices more likely to be software timed than their AI counterparts
Sample rate: Samples-per-second
  Channels not multiplexed
  Other hardware limitations prevent simultaneous update on most devices
  Look at spec sheet to determine maximum sample rate for N channels
  http://www.ni.com/pdf/manuals/370822c.pdf
New: Slew rate: maximum rate of change for output
New: Settling time: time required after a step change to achieve stable output.*
*Actually, ADCs have a settling time, but it’s not usually a limiting factor.

## Slide 66: Analog Output Hardware: Timing
http://www.planetanalog.com/author.asp?section_id=3062&doc_id=560483

## Slide 67: Affect of Slew Rate and Settling Time
Talk about the effects on output frequency

## Slide 68: Cool example I spent 90 minutes on

## Slide 69: Digital I/O Hardware
Wide variety of implementations—the following statements are most often true
Read the specs and docs!
Digital lines act like analog inputs or outputs EXCEPT:
  Lines typically bidirectional (may be configured to be either input or output)
  Lines are, of course, digital. Only read or generate two states called:
    High and Low
    On and Off
    One and Zero
    Etc.
* Not counting the “high-Z” state, meaning “not driven high or low”

## Slide 70: Digital I/O Hardware
Lines often grouped into “ports”
  Ports are typically 4 or 8 bits
  Sometimes only ports, not individual lines may be configured to be input or output
  Binary numbers may be written to ports in software
Same timing specifications apply to DIO lines:
  Hardware or software timing
    More likely to be software-only
  Sample speed for hardware timed acquisitions
  Slew rate → Rise time
  Settling time

## Slide 71: Using MAX DIO Test Panel
Follow along with this underwhelming example

## Slide 72: Understanding Logic Levels
Figure from https://learn.sparkfun.com/tutorials/logic-levels/ttl-logic-levels
Voltage | Description | TTL | LVTTL
Vcc | Supply voltage | 5.0 | 3.3
VOH | Logic-high output voltage minimum | 2.7 | 2.4
VIH | Logic-high input voltage minimum | 2.0 | 2.0
VIL | Logic-low input voltage maximum | 0.8 | 0.8
VOL | Logic-low output voltage maximum | 0.4 | 0.4
Levels are specified depending on whether they are inputs or outputs
This guarantees that outputs are always in the correct range for the inputs

## Slide 73: Many more logic families
Figure from http://www.interfacebus.com/voltage_LV_threshold.html
More logic levels
Most Devices

## Slide 74: How to solve logic level problems
Figures from Texas Instruments SN74LV4T125 Datasheet
Ignore it
  Some voltage levels are cross-compatible, e.g. most TTL and LVTTL
  However—if at all possible, check all output levels before connecting to the input
Use level-shifter circuit (voltage level translation)
  e.g. SN74LV4T125

## Slide 75: Analog to Digital
Use a comparator to help an analog signal conform to your digital standard
In
Out

## Slide 76: Reconfiguration Terminals
Many digital terminals mapped to multiple functions via software
  Allows more device functionality without increased connection count, cost and complexity
In NI terminology, these are “PFI” lines
Some lines will automatically map to default values depending on the software configuration
Some lines will need to be explicitly assigned in software
Make sure that you can map every function you intend to use to its own terminal simultaneously.
myDAQ User Manual, DIO section
PFI (Programmable Function Inputs) Lines

## Slide 77: Reconfiguration Terminals
Use NI-MAX to find function → terminal mappings

## Slide 78: Counter/Timers—Special DIO
Counter/Timers are a special class of digital I/O specifically designed for—you guessed it—counting and timing.
They can operate in several modes, as configured by software
  Measure or Generate
  Single Pulses or Pulse Trains

## Slide 79: Counter/Timers—Special DIO
Counter/Timers may be used to read Quadrature Encoders
(Linear or Rotary)
http://www.orientalmotor.com/technology/articles/servo-motor-glossary.html
Generates two pulse trains

## Slide 80: Counter/Timers—Special DIO
Quadrature Encoders
Count rising edges on A
http://www.robotoid.com/appnotes/images/quad-encoding-waveform.png
  If B is high,
  count down
  If B is low,
  count up

## Slide 81
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics

## Slide 82: Analog Input Software
Hardware Driver
  Provides the a software interface to the hardware
  For NI it’s NI-DAQmx
NI-DAQmx terminology:
  Channel: a single interface, e.g. one AI channel, one counter-timer, etc.
  Task: A acquisition or generation of:
    one or more data points
    on one or more channels
    of the same type
    and on the same device (usually).
  Scale: a calibration curve for Analog Input or Analog Outputs
Making your hardware work for you

## Slide 83: A Word about MAX
NI Measurement and Automation Explorer (MAX)
  Central location for investigating NI Hardware and Software
Allows for global (i.e. system-wide) channel and task creation
  Useful for fixed or infrequently-changed installations

## Slide 84: The Mantra
Create: Open a device for further access
  In other drivers: “open” or “initialize”
Configure: setup the task, including channel(s), range(s), timing, triggering, etc. for the task
Start: Begin the acquisition (or generation) process
Read/Write: read acquired data (or write data to be generated), typically repeatedly, until the task is complete
Stop: stop the acquisition but do not close the device and free the resources associated with it
Clear: okay, now close the device and free the resources
  In other drivers: “close” or “finalize”
Create, Configure, Start, Read/Write, Stop, Clear

## Slide 85: The Mantra
NI-DAQmx is “smart” and will attempt to do steps in the mantra that you “forget.”
If you don’t configure the task, defaults are chosen
If you don’t start the task before reading, it starts it when you attempt to perform the read
If you don’t stop the task before closing it, it stops it for you
For learners, especially, don’t “forget” to do any of the steps.
Leads to muddled thinking
May have unexpected results because the coder doesn’t understand what NI-DAQmx is doing under the hood.
A note about NI-DAQmx
Create, Configure, Start, Read/Write, Stop, Clear

## Slide 86: Software Timed Analog Input
Uses system clock
  Best timing ~ 1 ms period on modern, non-RT OSes
Subject to jitter due to processor multitasking
Must be used for process control systems
  Maybe you should consider an RT system?
NI Example: Voltage - SW timed Input.vi

## Slide 87: Software Timing: Jitter
Do Voltage – SW-Timed – Jitter.vi example

## Slide 88: Finite Samples, Hardware Timed
To relieve jitter and free up the process, let the DAQ card handle the timing
We must now configure the channel timing
  Use DAQmx Timing.vi
  Specify:
    Timing type (Finite samples in this case)
    Frequency
    Number of samples
    Timing source*
*Timing source allows one to use an external clock rather than an internal clock. This is useful for clocking in data based on external events signals, e.g. an encoder pulse, a clock from another DAQ task (e.g. analog output) or an external instrument.

## Slide 89: Finite Samples, Hardware Timed
See also NI Example: Voltage – Finite Input.vi

## Slide 90: What’s really happening?*
1 Configuration Step
*This is still a simplification of the real architecture. Sorry.

## Slide 91: What’s really happening?*
1 Configuration Step
N Reads w/o CPU

## Slide 92: What’s really happening?*
1 Configuration Step
N Reads w/o CPU
N Reads w/ CPU at our leisure

## Slide 93: Why is Hardware Timing better than Software?
1 Configuration Step
N Reads w/ CPU
N Writes w/ CPU
N Reads w/ CPU
All must be on time!

## Slide 94: Continuous Sampling
Acquiring data forever
But how will I graduate?
Specify number of samples to read.
Defaults to “all available.”
See also NI Example: Voltage – Continuous Input.vi
Notes: Modify Finite into Continuous

## Slide 95: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 96: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 97: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 98: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 99: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 100: What’s really happening?
Circular buffer overflows
Notes: Modify Finite into Continuous

## Slide 101: Buffer Overflows

## Slide 102: Acquiring for Multiple Channels
Use one task to acquire data from multiple channel
  Always works for multiple AI or DI channels on the same device
  Sometimes works for multiple channels on different devices
    Consider carefully the time skew implications. Are the devices synchronized?
    E.g. cDAQ devices share a common clock
    PXI Devices can share a common clock
    USB devices will not have a common clock unless externally synchronized.

## Slide 103: Acquiring for Multiple Channels
Specify multiple channels in the “physical channels” control
  Use commas to separate, and colons to specify ranges
Using a loop to add multiple channels to a task
  Allows for different ranges per-channel

## Slide 104: Analog Output Software
Just like analog input, but write instead of read
Voltage – Software Timed Output.vi

## Slide 105: Analog Output Software
Continuous Hardware Timed, e.g. function generator
Voltage – Continuous Hardware Timed Output.vi
See also: NI Example “Voltage – Continuous Output.vi”

## Slide 106: Analog Output Software
Continuous Hardware Timed, with write updates
FirstHallSensorSimulator.vi
See also: NI Example “Voltage – Continuous Output.vi”

## Slide 107: AO: Regeneration Mode
In continuous output operations, an important option is the “regeneration mode”
Regeneration Allowed: AO will continuously loop through data already in the buffer if no updates are available.
  Default mode
  Good for function generators and other continuous output
Regeneration Not Allowed: AO will require output buffer data to be updated as often as it is emptied.
  Failure to do so results in a buffer underflow error
Access with DAQmx Write Property Node

## Slide 108: Analog Output Software
Only update when you need to
Simulated Hall Effect Sensor.vi

## Slide 109: Digital Input/Output
If you can do analog input, you can do digital input
If you can do analog output, you can do digital output
Look a the examples in the Example Finder
  Hardware Input/Output->DAQmx->Digital Input/Output
Many devices only support software-timed DIO (including myDAQ)
  In a bind, AI/AO can substitute for DI/DO
Software-timed Digital Output

## Slide 110: Counter/Timer Input
Measure TTL pulse
  Period
  Frequency
  Pulse Width
  Count edges
  Etc.
CounterTimer SW-Timed.vi
See NI Examples
See Polymorphic “DAQmx Create Channel.vi”

## Slide 111: Counter/Timer Output
Generate
  Single Pulse
  Pulse Train
NI Examples
  Hardware I/O -> DAQmx -> Counter/Timer->*
NI Examples
Counter Output – Continuous –no trigger.vi

## Slide 112: Triggering
Use triggering to control when data acquisition occurs
Controls when the sample clock is active
Why would I want to do that?
  Only get the data you need based on external conditions
  Synchronize

## Slide 113: Triggering – Common Types
Digital start trigger: start when a rising or falling edge of a TTL (0-5 V) signal is detected
Digital reference trigger: similar to digital start, but may acquire pre-trigger samples
Digital pause trigger: clock in data only while a digital signal is high/low
Analog* start trigger: start when a signal’s voltage rises above or falls below a certain level
Analog* reference trigger: similar to analog start, but with pre-trigger samples
Analog* pause trigger: clock in data only while an analog signal is above/below a threshold
*Most NI DAQ devices do not support analog triggering. A few do not support digital triggering. See http://digital.ni.com/public.nsf/allkb/E761255A5729CD2E862570310075A23A

## Slide 114: Triggering
 | Digital | Analog
Start Trigger |  | 
Reference Trigger (Include pretrigger samples) |  | 
Pause Trigger (Gate) |  | 

## Slide 115: Triggering

## Slide 116: Digital Start Triggering
Show Digital – Finite Input Example

## Slide 117: External Clocks
90% of the time we use the DAQ scan clock to clock input or output conversions
Sometimes we may want to synchronize each sample to another event. Examples:
  Synchronize data acquisition/generation to an external instrument
  Synchronize multiple tasks on a single device (e.g. AI and AO tasks)
  Synchronize data acquisition/generation to another DAQ device
  Acquire data by position rather than time by clocking with an encoder pulse
  Create crazy acquisition patterns by using counter-timers, gates, external circuitry, etc.

## Slide 118: Synchronizing to external clock
No different than normal hardware-timed I/O except for “Sample Clock Source” input to DAQmx Timing
Voltage – Continuous Input-ExtClk.vi

## Slide 119: Sharing clocks with multiple tasks
It is possible to internally route the clock line from one task to another
  On the same device
  On different devices within a CompactPCI or PXI chassis
  On different PCI/PCIe devices using the RTSI cable*
Sometimes you may need to manually route through PXI Trigger/RTSI lines using DAQmx Connect Terminals.vi
*http://digital.ni.com/public.nsf/allkb/A120195AAAA9222A86256C69007C8B27

## Slide 120: Sharing clocks with multiple tasks
Master Task: the one that generates the clock
Slave Task: a task that uses the same clock as the master
Always start the slave task(s) first!!
Synch AI-AO.llb

## Slide 121: The Power of Triggers and Clocks
Timing a I recently created using triggers and clocks

## Slide 122: Problems with Digital Triggers and Clocks
Wrong polarity
  Triggers look for rising or falling edges (or sometimes either)
Wrong voltage
  Digital logic operates on specific voltage levels
Glitches

## Slide 123: Problems with Digital Triggers and Clocks
Figure from http://www.interfacebus.com/voltage_LV_threshold.html
More logic levels

## Slide 124: Problems with Digital Triggers and Clocks
Figures from Texas Instruments SN74LV4T125 Datasheet
How to solve level problems
Ignore it
  Some voltage levels are cross-compatible, e.g. most TTL and LVTTL
  However—if at all possible, check all output levels before connecting to the input
Use level-shifter circuit (voltage level translation)
  e.g. SN74LV4T125

## Slide 125: Problems with Digital Triggers and Clocks
Glitches
Glitches occur when noise causes a signal to rise above and then fall below the threshold very quickly
Glitch

## Slide 126: Review
Create
Configure
Start
Read/Write
Stop
Clear

## Slide 127: About Floating Sources
Unreferenced sources may float at potentials different from the DAQ potential
  Unpowered devices
  Devices power by batteries
  Devices powered by isolated power supplies—Just because it’s plugged in doesn’t mean it’s grounded

## Slide 128: Bias resistors
Differential signals may float outside of the operating range of the ADC
  Floating source: powered by batteries or isolated supply
Bias resistors can hold the voltages in range by preventing the source from floating
  Make sure source is actually floating!
Value R should be such that:
Output impedance < R < Input impedance
Typically 10 kΩ – 100 kΩ
Further reading: http://www.ni.com/white-paper/3344/en/#toc3

## Slide 129
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics