## Slide 1: Getting Data in the Lab
LabVIEW, Data Acquisition,
and Instrument Control
Dr. Jason D. Sommerville
President, UPVI, LLC.
jason.sommerville@upvi.net
www.upvi.net
Copyright © 2015 UPVI, LLC. Redistribution is prohibited.

## Slide 2: Getting Data in the Lab
Part 2: Data Acquisition
Copyright © 2015 UPVI, LLC. Redistribution is prohibited.

## Slide 3
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics

## Slide 4: Data Acquisition
Data Acquisition or DAQ is the process of converting real-world measurements into digital computer data
  Signal measurement
  Signal digitization
  Moving the data to computer memory and/or disk
DAQ may also encompass signal generation
  Waveform generation
  Triggering
  Control systems
In 2015, if you’re writing data down by hand, you’re probably doing it wrong

## Slide 5: Types of DAQ: Signal Acquisition
Convert analog voltage signals into digital data
Single-shot or timed measurements
Analog Input
Digital Input
Counter/Timers
Use an analog-to-digital converter
Examples: temperature, sounds, voltage
Single-shot or timed
Single bit or bus
Example: proximity sensor, photointerrupter
Count events
Time events
Examples: encoders, pulse train timing, delay timing, delay generation
101110110100 101110100101 011001100101
110011000101 011100101010 110110001000
1 or 0
Buffer
555
4, 5, 6, …
10 ms
10 ms
Counter/Timer

## Slide 6: Types of DAQ: Generation
Convert digital data into analog waveforms
Single-shot or timed generation
Analog Output
Digital Output
Counter/Timers
Use a digital-to-analog converter
Examples: sound, system excitation, motor control
Single-shot or timed
Single bit or bus
Generating a slow trigger, controlling an LED, controlling a bus
Control timing
Examples: Pulse train generation, delay generation
1 or 0
Buffer
555
4, 5, 6, …
10 ms
10 ms
Counter/Timer
101110110100 101110100101 011001100101
110011000101 011100101010 110110001000

## Slide 7
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Transducer
Thermocouple

## Slide 8: Transducers
Convert real world data into something that can be measured by a data acquisition device
  Voltage
  Current
  Resistance

## Slide 9: Transducer Menagerie
Temperature | Pressure | Flow
Thermocouple Thermistor RTD IC Temp Sensor IR Sensor Fiber Bragg Grating (FBG) | Piezo-resistive strain gauge Capacitive Electromagnetic Piezoelectric Microphone FBG | Differential Pressure Pitot tube Calorimetric Turbine Electromagnetic Ultrasonic
Light | Position/Velocity | Strain, weight, etc
Photodiode Phototransistor Photomultiplier Single-photon avalanch diode CCD/CMOS array | Optical Encoders (A,L) LVDT (L) Resolver (A) Potentiometer (A) String Pot (L) Radar (L) Laser interferometry (L) | Strain gauge Load cell Piezoelectrics

## Slide 10: How to chose a transducer
What type of measurement do you need to make?
Range (25 – 300 ºC, 1 – 100 mN, etc.)
Accuracy (±1 ºC, ±0.1 mN, etc.)
  Repeatability
  Linearity—can be handled in software
  Hysteresis—is it the same going up as coming back down?
Timing
  Bandwidth (1 Hz, 100 kHz – 1 MHz, etc.)
  Response time
Output
  Analog or digital
  Analog: voltage, current, impedance, etc.
  What signal conditioning (or excitation) does it require?
  How will I connect it to my DAQ system?
Environmental isolation
  Voltage, temperature, noise, vacuum, water, etc.
Cost

## Slide 11: Example 1: Choosing a thermocouple
Thermocouples measure the difference in temperature between the hot junction and the cold junction
There are different types of thermocouples for different temperature ranges
More reading: http://www.ni.com/white-paper/4218/en/

## Slide 12: Example 1: Choosing a thermocouple
http://www.temprel.com/support/how-to-choose-thermocouple.aspx

## Slide 13
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Signal Conditioner
Amplifier, Low pass filter
Transducer
Thermocouple

## Slide 14: Signal Conditioning
Modify the transducer output so that it can be accurately measured by the DAQ system
Some basic signal conditioning is often included with some transducers and some DAQ systems
  X -> Voltage conversion
  Amplification
  Anti-aliasing filters
Amplification/Attenuation
  Small signals (typ. <100 mV) should be amplified to a range readable by the ADC
  Large signals (>5 – 10 V) voltage signals need to be attenuated
Filtering
  Noise: use low-pass or notch filters to remove known noise sources
  Anti-aliasing filter: a low-pass filter designed to prevent high-frequency components from being digitized and aliased by the DAQ system

## Slide 15: Signal Conditioning (cont.)
Voltage Isolation
  Allows signals at high potential differences to be safely connected to a DAQ system
Linearization
  Corrects for non-linearities in the transducer
  In modern systems this is usually done in software
Transconductance, etc.
  Convert current signals to voltage signals
  Other conversions possible as well (i.e. resistance to voltage, pulse-train to voltage, etc.)
Excitation
  Some sensors, e.g. RTDs and strain gauges, require a power source to excite them
  Not strictly signal conditioning

## Slide 16: Signal Conditioning—Hardware
Amplifiers and filters often build into DAQ hardware
Rack and rail systems
  NI SCXI
  Omega DRG or OM7
Custom solutions
  Build it yourself
NI SCXI
Omega DRG Module

## Slide 17
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ

## Slide 18: Analog Input Digitization
Analog signal: 3 Hz Sine
Sampling Period: 50 ms (20 Hz)
Digitized Samples
Sample Index | Time (s) | Voltage (V)
0 | 0.00 | 1.7101
1 | 0.05 | 4.80631
2 | 0.10 | 3.94005
3 | 0.15 | -0.174497
4 | 0.20 | -4.14519
5 | 0.25 | -4.69846
6 | 0.30 | -1.37819
7 | 0.35 | 3.07831
8 | 0.40 | 4.99695
9 | 0.45 | 2.79596
10 | 0.50 | -1.7101
11 | 0.55 | -4.80631
12 | 0.60 | -3.94005
13 | 0.65 | 0.174497
14 | 0.70 | 4.14519
15 | 0.75 | 4.69846
16 | 0.80 | 1.37819
17 | 0.85 | -3.07831
18 | 0.90 | -4.99695
19 | 0.95 | -2.79596

## Slide 19: DAQ: Analog Input
Central component is the Analog-to-Digital (ADC) converter
  Digitizes the analog signal and making a digital measurement available for storage in memory
  ADC must be supported by amplifiers, clocks, multiplexers, and other items
Typical DAQ-AI system has the following parameters:
  Differential vs. Single-ended inputs
  Voltage range and polarity (bipolar or unipolar)
  Gain settings
  ADC resolution
  Sampling speed and bandwidth
  Multi-channel multiplexing

## Slide 20: Analog Input
Differential Input (DI)

## Slide 21: Analog Input
Single-ended Input (SE)

## Slide 22: Analog Input: DI vs. SE

## Slide 23: Analog Input: DI vs. SE
Differential Input
Used with balanced signal lines
Less susceptible to noise
Less likelihood of ground loops
Lower channel count
Single-Ended
Used with unbalanced signals
Increased susceptibility to noise
Increased likelihood of ground loops
Higher channel count
Voltage difference must be within DAQ range
Neither side must float too far from DAQ ground
Differential Input Caveats

## Slide 24: Bias resistors
Differential signals may float outside of the operating range of the ADC
  Floating source: powered by batteries or isolated supply
Bias resistors can hold the voltages in range by preventing the source from floating
  Make sure source is actually floating!
Value R should be such that:
Output impedance < R < Input impedance
Typically 10 kΩ – 100 kΩ
Further reading: http://www.ni.com/white-paper/3344/en/#toc3

## Slide 25: Analog Input
Analog signal: 3 Hz Sine
Sampling Period: 50 ms (Rate is 20 Hz)
Digitized Samples
Bi-polar, Range: ±5 V
Resolution: 16 bits across the range

## Slide 26: ADC Range and Polarity
Range: the maximum and minimum voltage that may be read
Unipolar: only read signals >= 0 V (or only <= 0, sometimes)
Bipolar: may read positive and negative signals
Bi-polar, Range: ±5 V
Most Common Ranges:
±5 V, ±10 V

## Slide 27: Gain Settings
Configurable gain settings allow ADC to make more use of available ADC resolution
Note that noise is also amplified—better to amplify at source, if possible
Signal here must be (e.g.) ±5 V to make full use of ADC resolution
Gain is software-configurable

## Slide 28: ADC Resolution
The resolution with which a signal can be measured, specified in bits
Resolution in volts is given by ADC Range, Gain setting, and ADC resolution
Note that sometimes the specified “range” will be the product of the smallest gain and range, or all range-gain products will be specified.
Most Common Resolutions:
8, 12, 16, 24
Example:

## Slide 29: ADC Resolution: 3 bits

## Slide 30: ADC Resolution: 4 bits

## Slide 31: ADC Resolution: 8 bits

## Slide 32: Sampling Speed
The maximum rate that the hardware can make ADC conversions
In samples per second
  E.g. 250 S/s, 350 kS/S, 10 MS/s
Hardware or software controlled?
  Low-end hardware must be clocked by software—very unreliable timing on modern OSes
  Lab-quality software will use hardware timing
What about bandwidth?
  A system may be able to sample above its bandwidth
  Imagine putting a 10 Hz low-pass filter ahead of your 10 MS/s ADC
  Bandwidth is set by the analog characteristics of the electronics
  Bandwidth may be affected by your setup, esp. cabling.
  Usually not a concern until you get to 10s – 100s of MHz

## Slide 33: Undersampling
Undersampled data causes aliasing
The alias looks like a lower frequency signal
Analog signal: 35 Hz Sine
Sampling Period: 50 ms (20 Hz)
Digitized Samples
Alias Period: 200 ms (5 Hz)

## Slide 34: Nyquist-Shannon Sampling Theory
“The sampling frequency must be greater than or equal to twice the bandwidth of the signal sampled to determine the frequency content of the sample.”
If bandwidth starts at 0 Hz (usual), BW is simply the maximum frequency component in our sample.
If you need to determine the non-sinusoidal shape of a signal repeating at X Hz

## Slide 35: Anti-alias Filters
If you know you have frequency components which are not interesting, use an anti-alias filter.
Prevents aliasing by preventing the signal!

## Slide 36: Analog Input: Multiplexing
Scan Clock: Acquire one data point from every channel
Sample Clock: Acquire one data point from any channel

## Slide 37: Analog Input: Time Skew
Scan Clock
10 Hz
Sample Clock
100 Hz

## Slide 38: Dealing with Time Skew
Live with it
  Does it really matter for your data?
  Make sure you understand the effect it will have on your data
  By far the most common solution
Sample-and-hold circuitry
  An analog circuit which captures all channels’ voltages simultaneously
  Allows ADC and MUX to cycle through sampled voltages at scan clock rate
Multiple ADCs

## Slide 39: Analog Input: Full Disclosure
Gain circuitry actually comes after the MUXs in most designs.
This explains the tradeoff between DI/SE and channel count

## Slide 40: Some Examples
NI CompactDAQ Platform
USB connectivity or embedded controller
Modules support AI, AO, DIO Counter/timers, some signal conditioning
$800+
NI USB-6008
USB connectivity
12-bit AI 8 SE/4 DI
10 kS/s
2 AO
$189
Meas. Computing PCI-2513
PCI card
16-bit AI 16 SE/8 DI
1 MS/s
DIO, counter/timers
$909
NI PXIe Platform
Embedded controller or PCIe bridge
Modules support everything
$2000+

## Slide 41: How to choose DAQ hardware
What type of signals do you need to acquire/generate?
How many channels of each type?
After external conditioning, what will the polarity and range of your signal be?
DI/SE?
What resolution do you need across that range?
What bandwidth do you need?
At what rate do you need to acquire/generate the data?
  Note that bandwidth ≠ sampling frequency, but related
  How precise are you timing requirements? Can you use software timing?
  How much time-skew is acceptable?
Memory depth if signals cannot be transferred to system memory
Connectors
Form factor
Software interface
Front panel
Cost

## Slide 42: Put some example systems here!

## Slide 43
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics

## Slide 44: Analog Input Software
Hardware Driver
  Provides the a software interface to the hardware
  For NI it’s NI-DAQmx
NI-DAQmx terminology:
  Channel: a single interface, e.g. one AI channel, one counter-timer, etc.
  Task: A acquisition or generation of one or more data points on one or more channels of the same type.
Making your hardware work for you

## Slide 45: A Word about MAX
NI Measurement and Automation Explorer (MAX)
  Central location for investigating NI Hardware and Software
Allows for global (i.e. system-wide) channel and task creation
  Useful for fixed or infrequently-changed installations

## Slide 46: The Mantra
Open: Open a device for further access
Configure: setup the task, including channel(s), range(s), timing, triggering, etc. for the task
Start: Begin the acquisition (or generation) process
Read/Write: read acquired data (or write data to be generated), typically repeatedly, until the task is complete
Stop: stop the acquisition but do not close the device and free the resources associated with it
Close: okay, now close the device and free the resources
Open, Configure, Start, Read/Write, Stop, Close

## Slide 47: The Mantra
NI-DAQmx is “smart” and will attempt to do steps in the mantra that you “forget.”
If you don’t configure the task, defaults are chosen
If you don’t start the task before reading, it starts it when you attempt to perform the read
If you don’t stop the task before closing it, it stops it for you
For learners, especially, don’t “forget” to do any of the steps.
Leads to muddled thinking
May have unexpected results because the coder doesn’t understand what NI-DAQmx is doing under the hood.
A note about NI-DAQmx
Open, Configure, Start, Read/Write, Stop, Close

## Slide 48: Software Timed Analog Input
Uses system clock
  Best timing ~ 1 ms period on modern, non-RT OSes
Subject to jitter due to processor multitasking
Must be used for process control systems
  Maybe you should consider an RT system?
NI Example: Voltage - SW timed Input.vi

## Slide 49: Software Timing: Jitter
Show jitter example

## Slide 50: Finite Samples, Hardware Timed
To relieve jitter and free up the process, let the DAQ card handle the timing
We must now configure the channel timing
  Use DAQmx Timing.vi
  Specify:
    Timing type (Finite samples in this case)
    Frequency
    Number of samples
    Timing source*
*Timing source allows one to use an external clock rather than an internal clock. This is useful for clocking in data based on external events signals, e.g. an encoder pulse, a clock from another DAQ task (e.g. analog output) or an external instrument.

## Slide 51: Finite Samples, Hardware Timed

## Slide 52: What’s really happening?*
1 Configuration Step
*This is still a simplification of the real architecture. Sorry.

## Slide 53: What’s really happening?*
1 Configuration Step
N Reads w/o CPU

## Slide 54: What’s really happening?*
1 Configuration Step
N Reads w/o CPU
N Reads w/ CPU at our leisure

## Slide 55: Why is DMA better than PIO?
1 Configuration Step
N Reads w/ CPU
N Writes w/ CPU
N Reads w/ CPU
All must be on time!

## Slide 56: Continuous Sampling
Acquiring data forever
But how will I graduate?
Specify number of samples to read.
Defaults to “all available.”
Notes: Modify Finite into Continuous

## Slide 57: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 58: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 59: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 60: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 61: What’s really happening?
Circular Buffering
Notes: Modify Finite into Continuous

## Slide 62: What’s really happening?
Circular buffer overflows
Notes: Modify Finite into Continuous

## Slide 63: Buffer Overflows

## Slide 64: Triggering
Use triggering to delay the start of acquisition until a certain event occurs
Common Types:
  Digital start trigger: start when a rising or falling edge of a TTL (0-5 V) signal is detected
  Analog start trigger: start when a signal’s voltage rises above or falls below a certain level
  Digital reference trigger: similar to digital start, but may acquire pre-trigger samples
  Analog reference trigger: similar to analog start, but with pre-trigger samples

## Slide 65: Triggering

## Slide 66: Digital Start Triggering
Show Digital – Finite Input Example

## Slide 67: Problems with Digital Triggers
Wrong polarity
  Triggers look for rising or falling edges (or sometimes either)
Wrong voltage
  Digital logic operates on specific voltage levels
Glitches

## Slide 68: Problems with Digital Triggers
Figure from https://learn.sparkfun.com/tutorials/logic-levels/ttl-logic-levels
Voltage | Description | TTL | LVTTL
Vcc | Supply voltage | 5.0 | 3.3
VOH | Logic-high output voltage minimum | 2.7 | 2.4
VIH | Logic-high input voltage minimum | 2.0 | 2.0
VIL | Logic-low input voltage maximum | 0.8 | 0.8
VOL | Logic-low output voltage maximum | 0.4 | 0.4
TTL Logic Levels

## Slide 69: Problems with Digital Triggers
Figure from http://www.interfacebus.com/voltage_LV_threshold.html
More logic levels

## Slide 70: Problems with Digital Triggers
Figures from Texas Instruments SN74LV4T125 Datasheet
How to solve level problems
Ignore it
  Some voltage levels are cross-compatible, e.g. most TTL and LVTTL
  However—if at all possible, check all output levels before connecting to the input
Use level-shifter circuit (voltage level translation)
  e.g. SN74LV4T125

## Slide 71: Problems with Digital Triggers
Figure from http://www.interfacebus.com/voltage_LV_threshold.html
Glitches

## Slide 72: Review
Open
Configure
Start
Read/Write
Stop
Close

## Slide 73
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics
Phenomenon
Cooling of Tea
Transducer
Thermocouple
Signal Conditioner
Amplifier, Low pass filter
DAQ Hardware
NI CompactDAQ
Computer
Brought to you by IT
Analysis
Curve fitting
Plotting
DAQ
Analog Input
Storing Data
HDF5 Files
Thesis/
Dissertation
Graduate!
Software Basics